# Eye Motion Tracking in the CAVE

Kevin Richey

Computer Science Department

Virginia Tech

Blacksburg, VA    24060

krichey@rrinc.com

Friday 18 December 1998

# Table of Contents

| 1. Introduction | 1 |
| --- | --- |
| 2. Eye Tracker Set Up and Calibration in the CAVE | 1 |
| 2.1 Eye Tracker Hardware Setup | 1 |
| 2.2 Procedures to Power On the Eye Tracker | 2 |
| 2.3 Data Acquisition Software Configuration | 3 |
| 2.4 Eye Tracker CAVE Calibration with a User | 3 |
| Table 1: Values for Fixation Points | 5 |
| 4. Types of Eye Motion Data Collected in the CAVE | 5 |
| 5. Programmer’s Guide | 6 |
| 5.1 Programming the Eye Tracker in Performer Applications | 6 |
| 5.2 Programming the Eye Tracker in Non-Performer Applications | 6 |
| 5.3 LineOfSight Class | 7 |
| 5.4 EyeTracker Class | 8 |
| 6. Unresolved Issues | 10 |
| 7. References | 11 |
| Appendix A: EyeCal Source Code | 12 |
| eyecal.cc | 12 |
| eyestuff.h | 13 |
| eyestuff.cc | 13 |
| Appendix B: LineOfSight Class Source Code | 16 |
| los.h | 16 |
| los.cc | 17 |
| Appendix C: EyeTracker Class Source Code | 20 |
| eyetracker.h | 20 |
| eyetracker.cc | 21 |

# Abstract

This document contains technical information about how to use the ISCAN Eye Motion Tracking System in the Virginia Tech CAVE.    The eye tracker must be calibrated to each person who wears the headset.    EyeCal is a CAVE program that easily facilitates the calibration process.    Programmers must modify their CAVE C++ Performer applications to make use of the eye tracker classes.    The LineOfSight class, derived from the Performer class pfDCS, translates the eye motion data from the eye tracker into a three-dimensional line segment that is used for intersection testing with objects in the virtual environment.    The results of the intersection testing are written to a log file, which can later be analyzed.    At a lower level, the EyeTracker class provides an interface for reading the eye motion data from the eye tracker computer.    All source code is included in the appendices.

# 1. Introduction

The eye motion tracker in the Virginia Tech Cave Automated Virtual Environment (CAVE) collects three-dimensional measurements of a user’s eye inside the CAVE’s virtual environment.    These measurements are performed by calculating the direction of a line called the line-of-sight, which originates from the user’s eye and extends into the environment in the direction of the pupil’s gaze.    Intersections between the line-of-sight and visible objects in the virtual world are computed, and the results are stored to a log file on disk.

There are three general steps to using the eye motion tracker in the CAVE.    The first is to add functionality to a new or existing CAVE application’s source code to accommodate for the eye tracker.    C++ classes have been written to make this step easy for programs that use the Performer library.    Software that does not use Performer (such as OpenGL applications) must be programmed with a custom line of sight intersection testing routine.    This step is documented in the Programmer’s Guide section.

The next step is to set up the eye tracker system and calibrate it with a user who will wear the headset.    This step is discussed in the section titled Eye Tracker Setup and Calibration in the CAVE.    Everyone has slight physiological differences in their eye.    Calibration will make corrections to accommodate for these differences, so it must be performed every time a different person wears the eye tracker headset.

The last step is to perform the eye tracking session in the CAVE and collect data about the user’s eye movements, which are stored to a log file in the CAVE computer.    The exact nature of the eye tracking session depends on the CAVE application and how the investigator wishes to use the data, so this step is beyond the scope of this document.    Section 4: Types of Eye Motion Data Collected in the CAVE, describes the format of the log file and how the data can be interpreted.

As of this writing, there are a few issues which have not been resolved.    Each issue is explained in detail in section 6: Unresolved Issues.    One issue in particular is the incomplete EyeTracker C++ class, which reads the data from the eye tracker computer via serial cable.    This software must be finished before the eye tracker is usable in the CAVE.

# 2. Eye Tracker Set Up and Calibration in the CAVE

Setting up the eye motion tracker for use in the CAVE requires several steps which must be performed carefully in order for the eye tracker to function properly.    The eye tracker system, including the computer, camera boxes, and headset, is a delicate device which must be treated with care.    Here is a summary of the setup process.    Each step is described in detail in the following sections.

    - Set up the eye tracker hardware.

    - Power on the eye tracker.

    - Configure the Data Acquisition software on the eye tracker computer.

    - Calibrate the eye tracker to the user’s eye.

## 2.1 Eye Tracker Hardware Setup

This section explains the hardware setup for the eye motion tracking system.    Refer to figure 1 for a line drawing of the setup.    Before proceeding, make sure all devices are turned off and unplugged.    The eye camera black box is turned off when the switch is moved to the left.

    - Set up the computer (attach monitor, keyboard, and mouse).

    - Attach a video cable from the SONY Eye Monitor to the “Eye Video Out” port on the computer.

    - Attach a video cable from the “Eye Video In” port of the computer to the back of the eye camera black box.

    - Attach the eye camera cable from the headset to the front of the eye camera black box.

    - Plug the serial cable from the CAVE into the serial port on the computer.

    - Attach the white power box to the back of the eye camera, and plug the power cord into an outlet.

    - Plug the computer, monitor, and eye monitor power cords into electrical outlets.

Figure 1: Eye Tracker Hardware Setup

The eye tracking system also includes a scene camera (which looks out, in the same direction as the user is looking), and a head motion tracker.    These two items are not used in the CAVE because the scene camera does not provide any useful functionality, and the head tracker will interfere with the CAVE’s own head tracking sensor.

## 2.2 Procedures to Power On the Eye Tracker

Once the hardware has been assembled, it is now ready to be turned on.    The steps for powering on the equipment and starting the software must be followed carefully.    Improper use could result in software difficulties or physical damage to the hardware.

          - Turn on the computer, display, and the SONY eye monitor.

          - Make certain that the eye tracker headset is plugged into the eye camera black box.

CAUTION: Never power on the eye camera when it is not connected to the headset cable!    Doing so will damage the camera!

          - Turn on the power to the eye camera black box.

          - When DOS has displayed the C: prompt, go to the KERMIT directory.    Enter: cd \kermit

          - Run the kermit software: msk315.exe

          - Set the active port to COM2:    set port com2

          - Exit kermit: exit

          - Change to the ISCAN directory: cd \iscan

          - Run the eye tracker data acquisition software: daq3_50c.exe

At this point you will be running the eye motion tracking software on the computer, and the eye monitor will be displaying the image seen by the eye camera.    If you wear the headset, you should see an image of your eye in the monitor.

Steps four through seven need a little explanation.    Due to a conflict in the computer’s serial port, it is necessary to specifically set the active port to COM2.    Without these steps, the data acquisition software can lock up when it tries to send data through the serial port.

## 2.3 Data Acquisition Software Configuration

Before the data acquisition software can properly relay data from the eye tracker to the CAVE, it must be configured.    Normally the this section can be skipped because the computer remembers the configuration settings, but these steps must be performed the first time the eye tracker computer is set up.    It may also be necessary to check these settings for correctness if the eye tracker does not function correctly.

Configure the serial communications settings:

          - Access the “Config” menu:    Alt-C

          - Select the “Serial Data Output” option.

          - Set “Output Status” to On.

          - Set “Output Port” to COM2.

          - Set “Baud Rate” to 9600.

          - Close the “Serial Data Output” window: ESC key.

Adjust the pupil and corneal reflection thresholds:

          - If the “Threshold” is turned off, press the “T” key to turn it on.    A dot will appear in the circle.

          - If the “Pupil X-Hair” is turned off, press the “X” key to turn it on.

          - If the “C.R. X-Hair” is turned off, press the “H” key to turn it on.

          - If the “C.R. Limit” is turned off, press the “I” key to turn it on.

          - Press “F11” to highlight the pupil threshold meter.

          - While wearing the headset, use the up and down arrow keys to adjust the threshold until the image of your pupil in the eye monitor is filled with a solid white circle.

          - press “F12” to highlight the C.R. (corneal reflection) meter.

          - While wearing the headset, use the up and down arrow keys to adjust the threshold until the image of the small highlight on your cornea is a solid circle and the cross hair is tracking it accurately.

Set the serial output parameters:

          - Press the “O” key repeatedly to cycle through the options window (bottom-left of screen) until the “Serial Output Control” window appears.

          - Set the rows of the first column of output parameters to be “.Pupil.H1.”, “.Pupil.V1.”, and “.Pupil.D1.” respectively by pressing the S key to select a row and the P key to cycle through the parameters.    The second column of parameters should be blank.

## 2.4 Eye Tracker CAVE Calibration with a User

The eye motion tracker must be calibrated to the user’s eye.    Calibration is the process of mapping an individual’s physical eye movements to some known range of measurements.    In the CAVE, the eye’s motion is measured in degrees of rotation. Every person’s eyes have subtle physiological differences in size and shape.    Calibrating the eye tracker will adjust the software to correct for these differences.    This section explains how to use a simple CAVE application called EyeCal along with the data acquisition software for calibrating the eye motion tracker.

EyeCal is a simple CAVE application designed to assist with calibrating and testing the eye motion tracker.    On each physical wall of the CAVE is drawn a set of nine yellow points, arranged in a three by three grid and separated one foot apart.    These points, called fixation points, are used as visual references during calibration.    Pairs of red points are projected onto the CAVE’s floor, four feet away from each wall.    These points are used as markers for placing the user’s feet during calibration.    A virtual silver ball is drawn five feet in front of the user.    The ball moves with the direction of the head tracker, so it will appear to follow the direction in which the viewer is looking.    When the ball comes into contact with a wall or the floor, it will stick to the surface rather than pass through it, so it is always visible.

The eye calibration process maps the user’s physical eye movements to a range of known values, so the eye tracker can translate the motion of the eye into usable data.    Degrees are used to measure the angle of the pupil’s direction of gaze.    The yellow fixation points are set to be fourteen degrees apart when viewed from a distance of four feet.

To run the EyeCal software, type “./eyecal” from the directory where the software has been installed on the CAVE computer and press Enter.    You will first be given this prompt:

Eye level (feet):

Enter the vertical distance, in feet, from the floor to the user’s eyes.    Subtracting two or three inches from the user’s height is a good approximation.    You may use a decimal point to indicate fractions of a foot.    For example, to specify five and one half feet, you would enter:

Eye level (feet): 5.5

EyeCal will use this number to adjust the positions of the yellow fixation points so that the middle row will be exactly on eye level with the user.

After the eye level has been entered, the calibration room will be displayed in the CAVE.    If the center row of fixation points is above or below the user’s eye level, restart the program and try a new value for the eye level that corrects for the difference.

The last adjustment that you need to make is the silver ball’s alignment with the head tracker.    The head motion tracker may not be perfectly aligned with the user’s head, thus the silver ball may not be positioned directly in front of the user’s vision.    To correct any misalignment, have the user place both feet on a pair of red spots on the floor, face the nearest wall, and stare directly ahead at the center fixation point.    If the silver ball does not overlap the fixation point, then it must be adjusted.    The ball can be moved by pushing the small thumb joystick built into the CAVE’s wand.    Move the ball until it covers the center fixation point, and press the leftmost button on the wand.    The person wearing the eye tracker headset should be careful to keep his or her head steady and aimed directly at the fixation point.    You may use the joystick to adjust the ball’s alignment as much as you need to. To reset the ball to its original position, press the middle button.    [Note: EyeCal should be modified to output the correction angles to the screen, so that they may be reused in other eye tracking CAVE applications.]

Once EyeCal is adjusted, you may proceed with the calibration process.    This process will require a user to wear the eye tracking headset and step into the CAVE while a second person, the administrator, works with the eye tracker computer.    The first step is for the administrator to program the data acquisition software for the fixation points.    This is done through the “Calibration Control” window, which as displayed by pressing the “O” key to cycle through the options window (in the bottom-left of the screen).    This window shows a small rectangle with five dots plotted in a “+” formation.    These five points are virtual fixation points which will be mapped to the real fixation points in the CAVE.    Pressing the “S” key will cycle through each of the virtual fixation points by circling them one at a time.    The center point will be visited twice, once for horizontal angles and once for vertical angles, so there are actually six points that must be programmed.    At each point, the administrator must enter the appropriate angle into the “Values” box by pressing the “V” key, typing the number, and pressing “Enter.”    The data acquisition software will remember the numbers from one session to the next, so setting these values may not be necessary, but it is a good idea to verify them each time you use the eye tracker.    The values to be entered are listed in Table 1.

### Table 1: Values for Fixation Points

Fixation Point                      Value

----------------------------------

First center point                      0

Left point                                  -14

Right point                                +14

Second center point                    0

Top point                                    -14

Bottom point                              +14

The calibration process is now ready to begin.    The user wearing the eye tracker headset in the CAVE should place his or her feet on a pair of red spots on the floor and face the nearest wall.    The administrator can then follow these steps to calibrate the eye tracker with the user’s eye:

          - Press the “S” key repeatedly until the “Pupil D1” parameter is selected in the “Calibration Control” window.

          - Press the “P” key to calibrate the user’s pupil diameter.    The “Pixels” value in the calibration control window will update with a new pupil diameter value.

          - Perform following three steps (a through c) exactly six times, once for each fixation point:

            - Press the “S” key to select the next calibration point.

            - Instruct the user to stare at the yellow fixation point corresponding to the circled virtual point in the calibration control window.    The user should only move his or her eyes, leaving the head pointing straight ahead.    (The silver ball should not move from the middle point.)

            - Press the “P” key to calibrate that point.    The “Pixels” box will update with a new value.

# 4. Types of Eye Motion Data Collected in the CAVE

Six variables are collected and written to a log file during an eye tracking session in the CAVE: a time stamp, the user’s location within the navigable world (LOC), the direction of the user’s line-of-sight (LOS), the user’s pupil diameter, the point of intersection between the LOS and the closest object in the virtual world, and a textual description of the intersected object.    The time stamp is taken from the computer’s internal clock.    The LOS is a three-dimensional vector that points in the direction the user is looking in the CAVE’s virtual space.    Ray tracing is used by the Performer library to determine if the LOS intersects any objects in the world.    If an intersection does occur, the point at which the LOS hits the nearest object is recorded.    The object’s textual name is also recorded, if it has one (names are assigned to objects by the programmer of the CAVE application).

Additional information can be interpreted from the parameters in the log file by performing post-processing calculations such as interpolation or statistical sampling.    For example, eye motion distance, velocity, and acceleration can be computed by interpolating the LOS vector over time.    Other calculationed measurements may include the mean number of distinct objects intersected per second, the frequency at which a user diverts his or her eyes between two or more different objects, and the average distance a user must move his or her eye between objects.    The pupil diameter can be used as a good indication of how much light is entering the eye, and how the light intensity changes.

# 5. Programmer’s Guide

## 5.1 Programming the Eye Tracker in Performer Applications

This section explains how programmers can integrate the eye tracker with CAVE applications that are written using the Performer C++ library.    The LineOfSight class is a Performer node object which, when added to a three-dimensional scene, will compute the intersection of the nearest object within the user’s view.    Using the eye tracker requires, at the very least, adding five lines of code to a CAVE application:

LineOfSight *los = new LineOfSight;

los->setWorld(world);

los->setDist(1000);

scene->addChild(los);

scene->setTravMask(PFTRAV_ISECT, 0xffffffff,

PFTRAV_SELF | PFTRAV_DESCEND | PFTRAV_IS_CACHE, PF_OR);

The first line creates a new LineOfSight object.    The constructor takes one optional integer parameter, which if non-zero will make the silver eye-tracking ball visible.    By default the parameter is zero, which makes the ball invisible.

The second line uses the setWorld() member function to tell the LineOfSight object where it should begin searching the object hierarchy for intersections.    The parameter (in this case “world”) is an object derived from pfNode that is the root node to all the objects of interest to the eye tracker.    For simple applications, this can be the root node to the entire scene.    For complex environments, however, it would be a good idea to use a subset of the entire world.    Intersection testing is an expensive operation, and it can hinder performance if the virtual environment has a large number of objects.    For this reason, the designer of the world should consider placing all objects that will be tested for eye intersections under a single node, and pass this node as the parameter to setWorld().

The LineOfSight class reads data from the eye tracker computer and creates a line that originates from the viewer’s eye and extends into the direction the eye is looking.    The length of the line is specified with the setDist() member function (line three).    Objects that cross this line are tested for intersection.    Any object whose distance from the viewer is greater than the length of the line segment is not tested.    Finding a good length may require some experimentation, but in general it should be long enough to intersect objects that are within the user’s field of view.    A length that is too great will unnecessarily reduce system performance, and a length that is too short will miss objects that could be important.    Other factors to keep in mind are the scale of the virtual environment and the number of objects.    A sophisticated eye tracking CAVE program could adjust the length as the user navigates through the world, depending on the user’s navigational location and the number of objects in the immediate vicinity.

The last two lines of code add the LineOfSight object to the virtual environment’s scene and specify the traversal flags, which are used by the rendering engine for intersection testing.

## 5.2 Programming the Eye Tracker in Non-Performer Applications

Applications that do not use the Performer library must implement their own line-of-sight intersection testing.    The EyeTracker class can be used to read eye movement data, but the programmer must implement the code to translate the angles of movement into three-dimensional vectors.

For example, the translation from eye data into three-dimensional line-of-sight vectors is performed by the LineOfSight class by finding the forward direction of the head tracker and rotating it by the angles of eye movement.    This vector is then used in the intersection testing.

The EyeTracker class provides the interface for reading data from the eye tracker.    The open() method must be called with the TTY device file as a string parameter.    The TTY device is the file attached to the CAVE’s serial port.    Contact the CAVE system administrator for the exact name.    The function will return true (non-zero) on success and false (zero) upon error.    Once open, data can be read from the eye tracker.    The getEyeParams() method returns the three eye motion parameters of horizontal angle, vertical angle, and pupil diameter.    A true value (non-zero) is returned on successful read, and false (zero) if data were not available.    The bytes returned from getEyeParams() can then be packed into integers and used to compute the viewer’s line-of-sight.    See the EyeTracker Class section below for more details about these and other methods in the EyeTracker class.

## 5.3 LineOfSight Class

The LineOfSight class is derived from pfDCS and is used to perform intersection testing on objects in the virtual environment with an imaginary line segment drawn from the viewer’s eye into the direction the viewer is looking.

LineOfSight(int draw_ball)

The class constructor.    The single parameter draw_ball should be set to one of: LineOfSight::SHOW_BALL or LineOfSight::HIDE_BALL, to indicate whether the silver ball should be displayed.    The default value of this parameter is HIDE_BALL.

void setWorld(pfNode *node)

Tells the LineOfSight object where to search the object hierarchy for intersections between the viewer’s line-of-sight and the objects in the scene.    The parameter node should be the root node to the hierarchy in which to search.

void setDist(float d)

Sets the length of the line segment that extends from the viewer’s eye out into the virtual world.    This line is used for computing intersections with the viewer’s line-of-sight and the objects in the scene.    Any object whose distance from the viewer is greater than the parameter d will not be checked for intersection.

void setLogMode(LogMode)

[Not yet implemented.]    Sets the format for output to the log file.

int startLog(char *filename, float log_freq, LogMode mode)

[Not yet implemented.]    Initiates the logging of intersections between the viewer’s line-of-sight and objects in the scene.    Filename specifies the name of the log file to create, log_req sets the frequency (number of times per second) at which data should be written to the log, and mode specifies the format of the logging.

int    writeLog(char *message)

Writes a textual message to a currently open log file.

void closeLog()

Closes the log file, so messages are no longer written to it.

int SHOW_BALL

This constant can be used as a parameter to the LineOfSight constructor to indicate that the silver eye-tracking ball should be visible.    The default is to make the ball invisible.

int HIDE_BALL

This constant can be used as a parameter to the LineOfSight constructor to indicate that the silver eye-tracking ball is invisible.    This is the default.

## 5.4 EyeTracker Class

The EyeTracker class reads raw data from the eye tracker (via serial line) and translates it into usable parameters for the CAVE.    Currently, this class is unfinished.    The serial communications has been implemented, and the EyeData structure can be read from the eye tracker.    But the getEyeParams() member function, which translates the raw eye data into useful numbers has not been finished.

### EyeData Structure

typedef unsigned char byte;

struct EyeData {

byte prefix_0;      // Prefix byte = 0x44

byte prefix_1;      // Prefix byte = 0x44

byte h1_lsb;          // Pupil horizontal angle - least significant byte

byte h1_msb;          // Pupil horizontal angle - most significant byte

byte v1_lsb;          // Pupil vertical angle - least significant byte

byte v1_msb;          // Pupil vertical angle - most significant byte

byte d1_lsb;          // Pupil diameter - least significant byte

byte d1_msb;          // Pupil diameter - most significant byte

};

This structure is used for collecting data from the serial port.    Each parameter is sixteen bits long (two bytes) broken into a high and low byte.    This structure is filled during a call to EyeTracker::read().

void activate(void)

Sets the EyeTracker object to active status.    When the status is active, data are read from the communications port.    When the state is not active, eye motion data are simulated through the joystick.    [Note: this should change to not use the joystick.    Otherwise it could interfere with CAVE programs that use the joystick for other purposes.]

void deactivate(void)

Sets the EyeTracker to inactive status, so data are not read from the serial port.

int isActive(void)

Returns true (non-zero) if the EyeTracker object is in active status.

void    setOffsets(float h, float v)

float getHorzOffset(void)

float getVertOffset(void)

void    calCenter(void)

These functions are used to make corrections in the pupil angle for inaccuracies in the head motion tracker.    Offset values can be specified that are added to the horizontal and vertical pupil angles.    The setOffsets() method explicitly sets these offset values.    Parameter h sets the horizontal angle offset and v sets the vertical angle offset.    Calling this method with v and h both set to zero will reset the offsets.    The current offset values can be accessed through getHorzOffset() and getVertOffset().    Another method of specifying the offsets is to assign the current pupil angles to the offsets.    This is done through a call to calCenter().    The EyeTracker class instance will read the current eye tracking parameter and assign them to the horizontal and vertical offsets.    This is the technique used by the LineOfSight class when calibrating with the CAVE wand.

int open(char *device)

Creates the input stream for reading from the communications port via device, the UNIX TTY device driver connected to the communications port.    This method returns true (non-zero) if the data stream was successfully opened, otherwise false(zero) is returned on error.

int isOpen(void)

Returns true (non-zero) if the data input stream is open for reading from the communications port.

int read(EyeData *data)

Reads a single chunk of eye motion data from the communications port, storing the information in the structure referenced by parameter data.    This method returns true (non-zero) if the data was successfully read, or false (zero) if an error occurred.

int getEyeParams(int *horz, int *vert, int *diam)

[Not yet implemented]    Fills the parameters with data read from the input stream.    Horz returns the horizontal pupil angle, vert returns the vertical pupil angle, and diam returns the pupil’s diameter.    This method returns true (non-zero) on a successful read and false (zero) if an error occurred or if the data was malformed.    An error does not necessarily indicate that there is a problem.    The best way to handle an error is simply to ignore it and try reading the eye parameters again.    This function has not been properly implemented.

# 6. Unresolved Issues

As of this writing, there are a few unresolved issues that need to be addressed before the eye tracker can be fully applied in the CAVE.

          - The getEyeParams() method in the EyeTracker class is incomplete.    This method should call the read() method to get the next batch of eye motion parameters from the eye tracker, in the form of an EyeData structure (see section 5.4).    The two prefix bytes (prefix_0 and prefix_1) should be checked for validity (they must equal 44 hex).    Each of the three eye motion parameters (horizontal angle, vertical angle, and pupil diameter) must be packed into integers from their two-byte pairs in the structure.    Care must be taken to ensure that the sign bit is carried into the upper sixteen bits of the integer when the eye parameters are negative.

          - The logging functions of the LineOfSight class are unimplemented.    Specifically, the methods setLogMode(), startLog(), writeLog(), and closeLog() are unfinished.    startLog() needs to open a file pointer for writing,    writeLog() should record the eye motion parameters to the file pointer, and closeLog() should close the file.    Calls to these methods must also be made in order for the data to be written during the intersection testing loop.

          - The eye tracker headset has a head motion sensor.    The CAVE, however, has its own head motion sensor, which may not be compatible with the sensor on the eye tracker headset.    The eye tracker’s sensor should be removed and the CAVE’s sensor should be made to transfer onto the eye tracker headset.    See a CAVE administrator for assistance.

# 7. References

ISCAN-DAQ    ISCAN, Inc., Raw Eye Movement Data Acquisition Software Version 3.10 Operating Instructions

# Appendix A: EyeCal Source Code

## eyecal.cc

#include "eyestuff.h"

#include "los.h"

static void createScene(pfChannel *chan, float eye_level, int argc, char **argv);

int main(int argc,char **argv)

{

pfInit();

pfCAVEConfig(&argc,argv,NULL);

pfConfig();

pfCAVEInitChannels();

float eye_level;

if (CAVEConfig->Simulator)

pfCAVEMasterChan()->getFStats()->setClass(

PFSTATS_ALL^PFSTATSHW_ENGFXPIPE_FILL, PFSTATS_ON);

else

pfCAVEMasterChan()->getFStats()->setClass(PFSTATS_ALL, PFSTATS_OFF);

printf("Eye level (feet): ");

scanf("%f", &eye_level);

eye_level -= 0.5;

createScene(pfCAVEMasterChan(),eye_level, argc,argv);

pfDCS          head_DCS, nav_DCS;

pfMatrix    head_mat;

while (!CAVEgetbutton(CAVE_ESCKEY)){

pfSync();

pfCAVEPreFrame();

pfFrame();

pfCAVEPostFrame();

}

pfCAVEHalt();

pfExit();

return 0;

}

static void createScene(pfChannel *chan, float eye_level, int ,char **)

{

pfScene                *scene;

pfGeoState          *gstate;

pfLightSource    *light;

float                    x, y, z;

pfVec4                  sphere_color;

pfGroup                *world;

scene    = new pfScene;

gstate = new pfGeoState;

gstate->setMode(PFSTATE_ENLIGHTING, PF_ON);

gstate->setMode(PFSTATE_CULLFACE, PFCF_OFF);

scene->setGState(gstate);

// Add a light source

light = new pfLightSource;

light->setColor(PFLT_DIFFUSE, 1.0, 1.0, 1.0);

light->setAmbient(0.3, 0.3, 0.3);

scene->addChild(light);

// Create the world

world = new pfGroup;

world->addChild(createWall(pfVec3(-5,-5,0), pfVec3(-5, 5,0), 10));

world->addChild(createWall(pfVec3(-5, 5,0), pfVec3( 5, 5,0), 10));

world->addChild(createWall(pfVec3( 5, 5,0), pfVec3( 5,-5,0), 10));

world->addChild(createFloor());

// Add 9 Yellow spheres on each wall

sphere_color.set(1.0, 1.0, 0.0, 0.0);

for(z=eye_level-1; z<=eye_level+1; z+=1){

for(y=-1; y<=1; y+=1){

world->addChild(createSphere(-5, y, z, 0.07, sphere_color));

world->addChild(createSphere( 5, y, z, 0.07, sphere_color));

}

for(x=-1; x<=1; x+=1)

world->addChild(createSphere(x, 5.0, z, 0.07, sphere_color));

}

// Add Red spheres on floor (foot markers)

sphere_color.set(1.0, 0.0, 0.0, 0.0);

world->addChild(createSphere(-1,    -0.5, 0.0, 0.08, sphere_color));

world->addChild(createSphere(-1,      0.5, 0.0, 0.08, sphere_color));

world->addChild(createSphere(-0.5, 1,      0.0, 0.08, sphere_color));

world->addChild(createSphere( 0.5, 1,      0.0, 0.08, sphere_color));

world->addChild(createSphere( 1,    -0.5, 0.0, 0.08, sphere_color));

world->addChild(createSphere( 1,      0.5, 0.0, 0.08, sphere_color));

LineOfSight *los = new LineOfSight(LineOfSight::SHOW_BALL);

los->setWorld(world);

los->setDist(4);

scene->addChild(los);

scene->setTravMask(PFTRAV_ISECT, 0xffffffff,

PFTRAV_SELF | PFTRAV_DESCEND | PFTRAV_IS_CACHE, PF_OR);

scene->addChild(world);

chan->setScene(scene);

}

## eyestuff.h

#ifndef __EYESTUFF_H

#define __EYESTUFF_H

#include <stdio.h>

#include <stdlib.h>

#include <math.h>

#include <Performer/pf/pfChannel.h>

#include <Performer/pf/pfLightSource.h>

#include <Performer/pf/pfTraverser.h>

#include <Performer/pf/pfGeode.h>

#include <Performer/pr/pfGeoSet.h>

#include <Performer/pf/pfDCS.h>

#include <Performer/pfutil.h>

#include <Performer/pfdu.h>

#include <pfcave.h>

pfGeode    *createWall(pfVec3 corner1, pfVec3 corner2, float height);

pfGeode    *createFloor(void);

pfGeoSet *createGeoSet(pfVec3 *verts, pfVec4 *colors);

pfGeode    *createSphere(float xc, float yc, float zc, float radius,pfVec4 color);

#endif

## eyestuff.cc

#include "eyestuff.h"

pfGeode *createWall(pfVec3 corner1, pfVec3 corner2, float height)

{

pfGeoSet *wall_gset;

pfVec3      *wall_verts;

pfVec4      *wall_colors;

pfGeode    *geode;

// Create blue wall

wall_verts = (pfVec3 *) pfMalloc(4 * sizeof(pfVec3),

fGetSharedArena());

wall_verts[0] = corner1;

wall_verts[1] = corner2;

wall_verts[2] = corner2 + pfVec3(0,0,height);

wall_verts[3] = corner1 + pfVec3(0,0,height);

/* wall_verts[0].set(-5.0, 5.0,    0.0);

* wall_verts[1].set( 5.0, 5.0,    0.0);

* wall_verts[2].set( 5.0, 5.0, 10.0);

* wall_verts[3].set(-5.0, 5.0, 10.0);

*/

wall_colors =(pfVec4 *) pfMalloc(4 * sizeof(pfVec4),

pfGetSharedArena());

wall_colors[0].set(0.1, 0.1, 1.0, 1.0);

wall_colors[1].set(0.1, 0.1, 1.0, 1.0);

wall_colors[2].set(0.1, 0.1, 1.0, 1.0);

wall_colors[3].set(0.1, 0.1, 1.0, 1.0);

wall_gset    = createGeoSet(wall_verts, wall_colors);

geode = new pfGeode;

geode->addGSet(wall_gset);

geode->setName("Wall");

return geode;

}

pfGeode *createFloor(void)

{

pfGeoSet *floor_gset;

pfVec3      *floor_verts;

pfVec4      *floor_colors;

pfGeode    *geode;

// Create white floor

floor_verts = (pfVec3 *) pfMalloc(4 * sizeof(pfVec3),

pfGetSharedArena());

floor_verts[0].set(-6.0, 5.0,    0.0);

floor_verts[1].set(-6.0,-5.0,    0.0);

floor_verts[2].set( 6.0,-5.0,    0.0);

floor_verts[3].set( 6.0, 5.0,    0.0);

floor_colors =(pfVec4 *) pfMalloc(4 * sizeof(pfVec4),

pfGetSharedArena());

floor_colors[0].set(1.0, 1.0, 1.0, 1.0);

floor_colors[1].set(1.0, 1.0, 1.0, 1.0);

floor_colors[2].set(1.0, 1.0, 1.0, 1.0);

floor_colors[3].set(1.0, 1.0, 1.0, 1.0);

floor_gset    = createGeoSet(floor_verts, floor_colors);

geode = new pfGeode;

geode->addGSet(floor_gset);

geode->setName("Floor");

return geode;

}

pfGeoSet *createGeoSet(pfVec3 *verts, pfVec4 *colors)

{

pfGeoSet *gset;

gset = new pfGeoSet;

gset->setPrimType(PFGS_QUADS);

gset->setNumPrims(1);

gset->setAttr(PFGS_COORD3, PFGS_PER_VERTEX, verts,    NULL);

gset->setAttr(PFGS_COLOR4, PFGS_OVERALL,        colors, NULL);

gset->setGState(new pfGeoState);

return gset;

}

pfGeode *createSphere(float xc, float yc, float zc, float radius,pfVec4 color)

{

pfGeoSet *gset;

pfVec4 *colorAttr;

pfMatrix matrix;

pfGeode *geode = new pfGeode;

gset = pfdNewSphere(32,pfGetSharedArena());

colorAttr = (pfVec4 *) pfMalloc(sizeof(pfVec4),pfGetSharedArena());

*colorAttr = color;

gset->setAttr(PFGS_COLOR4,PFGS_OVERALL,colorAttr,NULL);

matrix.makeScale(radius,radius,radius);

pfdXformGSet(gset,matrix);

matrix.makeTrans(xc, yc, zc);

pfdXformGSet(gset,matrix);

geode->addGSet(gset);

return geode;

}

# Appendix B: LineOfSight Class Source Code

## los.h

#ifndef _los_h_

#define _los_h_

#include <time.h>

#include <Performer/pf/pfDCS.h>

#include <Performer/pr/pfLinMath.h>

#include "eyetracker.h"

#define PI_OVER_2                    1.570796

#define MESSAGE_LENGTH          300

#define DEFAULT_LOG_FREQ      1.0

typedef enum LOG_MODE_ENUM {

LOG_MODE_OFF,

LOG_MODE_ISECT,

LOG_MODE_EYE_MOVE

} LogMode;

class LineOfSight : public pfDCS {

//* Stuff for computing the line-of sight

public:

LineOfSight(int draw_ball = LineOfSight::HIDE_BALL);

void setWorld(pfNode *node);

void setDist(float d);

void setLogMode(LogMode);

int    startLog(char*, float log_freq = DEFAULT_LOG_FREQ,

LogMode = LOG_MODE_ISECT);

int    writeLog(char*);

void closeLog();

virtual ~LineOfSight() {}

static int SHOW_BALL;

static int HIDE_BALL;

static unsigned int ISECT_MASK;

private:

EyeTracker eye;

pfVec3          position;

pfVec3          direction;

pfVec3          target;

float            distance;

FILE              *logfile;

LogMode        log_mode;

float            log_frequency;

float            last_log_time;

int active_;

pfNode *world_;

// Stuff needed for APP traversal of node

public:

virtual int app(pfTraverser*);

virtual int needsApp(void);

private:

int hitObject(void);

// Stuff required to make this a Performer class

public:

static void init(void);

static pfType* getClassType(void){ return classType_; };

private:

static pfType* classType_;

};

#endif

## los.cc

#include <pfcave.h>

#include <Performer/pfdu.h>

#include <Performer/pf/pfGeode.h>

#include <Performer/pf/pfTraverser.h>

#include "los.h"

/************************************************************************/

/* Performer type data

*/

pfType *LineOfSight::classType_ = NULL;

void LineOfSight::init(void)

{

if (classType_ == NULL) {

pfDCS::init();				/* Initialize the parent class */

/* Create the new pfType for this class, based on the parent class's type */

classType_ =    new pfType(pfDCS::getClassType(), "LineOfSight");

}

}

/************************************************************************/

int LineOfSight::SHOW_BALL = 1;

int LineOfSight::HIDE_BALL = 0;

unsigned int LineOfSight::ISECT_MASK = 0x0100;

LineOfSight::LineOfSight(int draw_ball)

{

pfGeoSet *ball;

pfGeode *geode;

setType(classType_);    /* set the Performer type of this instance */

active_ = 1;

logfile    = NULL;

log_mode = LOG_MODE_OFF;

eye.activate();

//* Get head tracking info

CAVEGetPosition(CAVE_LEFT_EYE_NAV, position.vec);

CAVEGetVector(CAVE_HEAD_FRONT_NAV, direction.vec);

//* Make ball float in your line of sight

target = (direction * 30.0f) + position;

setTrans(target[0], target[1], target[2]);

if(draw_ball){

ball            = pfdNewSphere(24, pfGetSharedArena());

geode          = new pfGeode;

geode->addGSet(ball);

addChild(geode);

setScale(0.1);

}

setTravMask(PFTRAV_ISECT, 0, PFTRAV_SELF|PFTRAV_DESCEND, PF_SET);

}

void LineOfSight::setWorld(pfNode *node)

{

world_ = node;

distance = 5;

}

void LineOfSight::setDist(float d)

{

distance = d;

}

void LineOfSight::setLogMode(LogMode mode)

{

log_mode = mode;

}

int LineOfSight::startLog(char *filename, float freq, LogMode mode)

{

if(filename) logfile    = fopen(filename, "w");

if(logfile){

log_mode            = mode;

log_frequency = freq;

last_log_time = 0;

}

return (logfile != NULL);

}

int LineOfSight::writeLog(char *str)

{

int      result = 0;

float seconds = (float)clock() / (float)CLK_TCK;

if(seconds - last_log_time > log_frequency){

time_t the_time = time(NULL);

last_log_time      = seconds;

result = fprintf(logfile, "%s\t%s\n", ctime(&the_time), str);

}

return result;

}

void LineOfSight::closeLog()

{

fclose(logfile);

logfile    = NULL;

log_mode = LOG_MODE_OFF;

}

int LineOfSight::app(pfTraverser *trav)

{

pfVec3      head_front;

pfVec3      head_right;

pfVec3      head_up;

pfMatrix head_matrix;

float        horz_eye_angle;

float        vert_eye_angle;

int            eye_diameter;

pfVec3      eye_normal;

if(active_){

// Check for button clicks

// Button 1 is the calibrate button

if(CAVEButtonChange(1) && CAVEBUTTON1) eye.calCenter();

// Button 2 is the clear button

if(CAVEButtonChange(2) && CAVEBUTTON2) eye.setOffsets(0,0);

// Update head tracking info

CAVEGetPosition(CAVE_LEFT_EYE_NAV, position.vec);

CAVEGetVector(CAVE_HEAD_FRONT_NAV, head_front.vec);

CAVEGetVector(CAVE_HEAD_UP_NAV,        head_up.vec);

CAVEGetVector(CAVE_HEAD_RIGHT_NAV, head_right.vec);

head_matrix.makeIdent();

head_matrix.setCol(0, head_right);

head_matrix.setCol(1, head_front);

head_matrix.setCol(2, head_up);

head_matrix.invertAff(head_matrix);

// Get eye movement data

eye.getEyeParams(&horz_eye_angle, &vert_eye_angle, &eye_diameter);

horz_eye_angle = PI_OVER_2 - horz_eye_angle;

vert_eye_angle = PI_OVER_2 - vert_eye_angle;

// Compute direction vector of eye, normal to head orientation

eye_normal.set(sin(vert_eye_angle) * cos(horz_eye_angle),

sin(vert_eye_angle) * sin(horz_eye_angle),

cos(vert_eye_angle));

// Compute direction of viewer's eye

pfVec3 direction = eye_normal * head_matrix;

target = (direction * 5.0f) + position;

// If there is no intersection, make ball float 30 units in front

hitObject();

// Move ball to be in your line of sight

setTrans(target[0], target[1], target[2]);

}

return pfDCS::app(trav);

}

int LineOfSight::needsApp(void)

{

return TRUE;

}

int LineOfSight::hitObject(void)

{

pfSegSet segset;

pfHit        **hits[32];

char          *name;

char          log_msg[MESSAGE_LENGTH];

segset.activeMask = 1;

segset.isectMask    = 0xFFFF;

segset.discFunc      = NULL;

segset.bound            = NULL;

segset.mode              = PFTRAV_IS_PRIM;

segset.segs[0].makePts(position, target);

if(world_->isect(&segset, hits)){

pfVec3      isectPoint;

pfMatrix xmat;

(*hits[0])->query(PFQHIT_POINT, isectPoint.vec);

(*hits[0])->query(PFQHIT_XFORM, (float*)xmat.mat);

(*hits[0])->query(PFQHIT_NAME,    &name);

target.xformPt(isectPoint, xmat);

// Record intersection data in logfile

if(log_mode == LOG_MODE_ISECT){

sprintf(log_msg, "Look-At %s, <%f,%f,%f>", name,

target[0], target[1], target[2]);

writeLog(log_msg);

}

return 1;

}

return 0;

}

# Appendix C: EyeTracker Class Source Code

## eyetracker.h

#ifndef __EYETRACKER_H

#define __EYETRACKER_H

#include <stdio.h>

#include <stdlib.h>

#include <unistd.h>

#include <fcntl.h>

#include <math.h>

#include <Performer/pf/pfDCS.h>

#include <Performer/pfutil.h>

#include <Performer/pfdu.h>

#include <pfcave.h>

#define PI                                3.14159

#define RADIAN_SCALE            0.4363323

#define EYE_DATA_PREFIX      0x44

#define DEFAULT_DIAMETER    5.0

typedef unsigned char byte;

struct EyeData {

byte prefix_0;      // Prefix byte = 0x44

byte prefix_1;      // Prefix byte = 0x44

byte d1_lsb;          // Pupil diameter - least significant byte

byte d1_msb;          // Pupil diameter - most significant byte

byte h1_lsb;          // Pupil horizontal angle - least significant byte

byte h1_msb;          // Pupil horizontal angle - most significant byte

byte v1_lsb;          // Pupil vertical angle - least significant byte

byte v1_msb;          // Pupil vertical angle - most significant byte

};

class EyeTracker {

public:

EyeTracker();

void    activate(void);

void    deactivate(void);

int      isActive(void);

void    setOffsets(float h, float v);

float getHorzOffset(void);

float getVertOffset(void);

void    calCenter(void);

int      open(char*);

int      isOpen(void);

int      read(EyeData*);

int      getEyeParams(float *horz, float *vert, int *diam);

private:

float x_scale;

float y_scale;

float horz_offset, vert_offset;

int      active;

int      eye_stream;

};

#endif

## eyetracker.cc

#include "eyetracker.h"

EyeTracker::EyeTracker()

{

horz_offset    = 0;

vert_offset    = 0;

active              = 0;

eye_stream      = -1;

}

void EyeTracker::activate(void)

{

active = 1;

}

void EyeTracker::deactivate(void)

{

active = 0;

}

int EyeTracker::isActive(void)

{

return active;

}

void EyeTracker::setOffsets(float h, float v)

{

horz_offset = h;

vert_offset = v;

}

float EyeTracker::getHorzOffset(void)

{

return horz_offset;

}

float EyeTracker::getVertOffset(void)

{

return vert_offset;

}

void EyeTracker::calCenter(void)

int diameter;

getEyeParams(&horz_offset, &vert_offset, &diameter);

}

int EyeTracker::open(char *file)

{

if(file) eye_stream = ::open(file, 0);

else eye_stream = -1;

return (eye_stream != -1);

}

int EyeTracker::isOpen(void)

{

return (eye_stream != -1);

}

int EyeTracker::read(EyeData *eye_stuff)

{

if(eye_stuff) return ::read(eye_stream, eye_stuff, sizeof(EyeData));

else return 0;

}

int EyeTracker::getEyeParams(float *horz_angle, float *vert_angle, int *diameter)

{

EyeData eye_stuff;

int          num_read;

int          horz_deg, vert_deg;

int          result;

result = 1;

// If the eye tracker is active, compute eye angles

if(active){

// If the stream is open, read data from stream

if(eye_stream){

num_read = this->read(&eye_stuff);

if(num_read == sizeof(EyeData)){

if(eye_stuff.prefix_0 == EYE_DATA_PREFIX &&

eye_stuff.prefix_1 == EYE_DATA_PREFIX){

*diameter      = (eye_stuff.d1_msb << 8) + eye_stuff.d1_lsb;

horz_deg        = (eye_stuff.h1_msb << 8) + eye_stuff.h1_lsb;

*horz_angle = (((float)horz_deg * PI) / 180.0) + horz_offset;

vert_deg        = (eye_stuff.v1_msb << 8) + eye_stuff.v1_lsb;

*vert_angle = (((float)vert_deg * PI) / 180.0) + vert_offset;

} else result = 0;

} else result = 0;

}

// If stream is not open, simulate with joystick

else {

*diameter      = DEFAULT_DIAMETER;

*horz_angle = CAVE_JOYSTICK_X * RADIAN_SCALE + horz_offset;

*vert_angle = CAVE_JOYSTICK_Y * RADIAN_SCALE + vert_offset;

}

}

// If eye tracker is not active, just look straight ahead

else {

*diameter      = DEFAULT_DIAMETER;

*horz_angle = horz_offset;

*vert_angle = vert_offset;

}

return result;

}